Edge AI Dementia Voice Detector

Audio Pipeline — Technical Documentation

audio_pipeline.py  |  test_pipeline.py

Team: Zeina Hammoud · Celine Sadaka · Raoul Saber

Target Hardware: Raspberry Pi 4  |  Validated by: Dr. Khaled Sidani (Resident Neurosurgeon, AUBMC)

|  | IMPORTANT — Clinical Disclaimer: This system produces screening indicators only and is NOT a medical diagnostic device. All outputs are intended to support early pre-assessment and clinician referral. Any screening result must be reviewed by a qualified healthcare professional before clinical action is taken. |
| --- | --- |

# 1. Overview

The audio pipeline is the first stage of the dementia voice screening system. It is responsible for capturing speech from a USB microphone connected to a Raspberry Pi 4, cleaning that audio, and producing a standardised WAV file that can be passed directly to the inference engine.

The pipeline was intentionally designed to preserve all temporal structure in the audio — including long pauses, hesitations, and silence between words. These are primary acoustic markers for early cognitive decline, and trimming them would destroy the signal the model depends on.

## 1.1  What this pipeline does

- Captures audio from a USB microphone at 16 kHz, mono, 32-bit float

- Applies a bandpass filter to remove sub-80 Hz rumble and above-8 kHz noise

- Runs spectral subtraction noise reduction using the first 0.5 s as a noise profile

- Normalises peak amplitude to -3 dB for consistent model input levels

- Saves both raw and cleaned WAV files for side-by-side quality validation

- Prints a full processing summary including SNR before/after, clipping status, and speech ratio

## 1.2  What this pipeline does NOT do

- Does not run inference or load any ML model — that is handled by a separate module

- Does not trim silence or pauses — VAD is implemented but disabled by default

- Does not extract features (MFCCs, pitch, etc.) — the model handles that internally

- Does not send audio to the cloud — all processing is fully offline on-device

# 2. File Structure

| File | Role |
| --- | --- |
| audio_pipeline.py | Core library — all signal processing logic and data classes. Import this in other modules. |
| test_pipeline.py | Standalone test script — records from mic, runs the pipeline, saves outputs, prints summary. |

## 2.1  audio_pipeline.py — Module Structure

| Section / Symbol | Type | Description |
| --- | --- | --- |
| PipelineStage | Enum | Identifies each processing stage: RAW, FILTERED, DENOISED, VAD_TRIMMED, NORMALIZED |
| PipelineConfig | Dataclass | All configuration knobs in one place. Instantiate and pass to run_pipeline(). |
| PipelineResult | Dataclass | Return type of run_pipeline(). Contains cleaned audio and all diagnostic fields. |
| MicDisconnectedError | Exception | Raised by capture_audio() if the USB mic disappears or is unavailable. |
| SilenceOnlyError | Exception | Raised by voice_activity_detection() if no speech is detected (VAD mode only). |
| list_audio_devices() | Function | Prints all ALSA/PortAudio input devices. Use to find the USB mic device index. |
| capture_audio() | Function | Records from the microphone. Returns float32 numpy array shape (N,). |
| bandpass_filter() | Function | 5th-order Butterworth filter, 80 Hz – 8 kHz, zero-phase. |
| spectral_subtraction() | Function | Noise reduction using FFT-based spectral subtraction with oversubtraction. |
| voice_activity_detection() | Function | Energy + ZCR-based VAD. Disabled by default. Available for diagnostic use. |
| normalize_audio() | Function | Peak normalisation to configurable target dB level. |
| check_clipping() | Function | Detects clipping/distortion in raw audio. Always runs before processing. |
| estimate_snr_db() | Function | Estimates SNR by comparing first-segment noise RMS to full-signal RMS. |
| run_pipeline() | Function | Orchestrates all stages. Main entry point for programmatic use. |
| save_wav() | Function | Saves float32 audio as 16-bit PCM WAV. |
| load_wav() | Function | Loads a WAV file and returns (float32 array, sample_rate). |

# 3. Signal Processing Chain

The pipeline runs four stages in sequence. Each stage is independently skippable via the skip_stages parameter in run_pipeline(), which is useful for debugging and isolation testing.

## Stage 1 — Bandpass Filter

Removes frequencies outside the speech range. Implemented as a 5th-order Butterworth filter using scipy's sosfiltfilt (zero-phase, no timing shift).

| Parameter | Default | Effect |
| --- | --- | --- |
| bp_low_hz | 80 Hz | Cut-off for low-frequency rumble (HVAC, handling noise) |
| bp_high_hz | 8000 Hz | Cut-off for high-frequency hiss (above speech formants) |
| bp_order | 5 | Filter steepness — higher order = sharper rolloff but more computation |

CPU cost on Pi 4: ~50 ms for 30 s of audio.

## Stage 2 — Spectral Subtraction Noise Reduction

Estimates the ambient noise spectrum from the first 0.5 s of the recording, then subtracts that estimate from every subsequent frame in the frequency domain. The original phase is preserved so speech timing is not altered.

Algorithm per frame:

- Compute FFT of windowed frame (1024-point Hanning window)

- Subtract: cleaned_magnitude = |signal| − alpha × |noise|

- Apply spectral floor: max(cleaned_magnitude, beta × |signal|) to prevent musical noise

- Reconstruct using original phase and compute IFFT

- Overlap-add with 75% overlap (hop = 256 samples = 16 ms)

| Parameter | Default | Effect |
| --- | --- | --- |
| noise_estimate_duration_s | 0.5 s | Length of initial silent segment used to estimate noise profile |
| n_fft | 1024 | FFT size — 64 ms window at 16 kHz. Larger = better frequency resolution |
| hop_length | 256 | Hop size — 16 ms. Smaller = smoother but more CPU |
| oversubtraction | 2.0 | Alpha: aggressiveness (1.0 = mild, 4.0 = aggressive). Increase for noisy rooms |
| spectral_floor | 0.02 | Beta: minimum residual. Prevents over-suppression and musical noise artefacts |

CPU cost on Pi 4: ~200–400 ms for 30 s of audio. Peak memory: ~3 MB for FFT workspace.

|  | Noise Profile Tip The first 0.5 s of every recording is used as the noise reference. For best results, ask the subject to remain silent for the first half-second before speaking. In noisy environments, increase noise_estimate_duration_s to 1.0 s or set oversubtraction to 3.0. |
| --- | --- |

## Stage 3 — Voice Activity Detection (Disabled by Default)

VAD is implemented but intentionally skipped by default. Long pauses, hesitations, and silence between words are clinically significant markers for early dementia, so trimming them would remove exactly the signal the model needs to detect.

VAD remains available for diagnostic use — for example, to measure the ratio of speech to silence in a recording, or to verify that the subject actually spoke. It uses energy (RMS in dB), zero-crossing rate, hangover logic, and median smoothing to produce a per-frame speech/silence mask.

| Parameter | Default | Description |
| --- | --- | --- |
| vad_energy_threshold_db | -35 dB | Energy threshold relative to peak. Frames below this are silence |
| vad_zcr_threshold | 0.30 | Zero-crossing rate upper bound — high ZCR frames are likely noise |
| vad_min_speech_ms | 200 ms | Minimum duration for a speech chunk to be kept |
| vad_padding_ms | 150 ms | Context padding added around detected speech segments |
| vad_hangover_frames | 8 | Number of frames to keep after energy drops (catches word endings) |

## Stage 4 — Normalisation

Peak normalises the cleaned audio to a configurable dB level (default -3 dBFS). This ensures the model receives a consistent input amplitude regardless of how close the subject was to the microphone or how loud the environment was.

RMS normalisation is not applied by default, as it would alter the dynamic range that the model may rely on for prosody and energy-based features. If the teammate's model requires RMS-normalised input, set target_rms_db in PipelineConfig.

CPU cost on Pi 4: < 10 ms for 30 s of audio.

# 4. Configuration Reference — PipelineConfig

All pipeline parameters are controlled through a single PipelineConfig dataclass. Instantiate it with keyword overrides — anything not specified uses the default value. Never hardcode values in the processing functions.

## 4.1  Audio Capture Parameters

| Parameter | Type | Default | Description |
| --- | --- | --- | --- |
| sample_rate | int | 16000 | Sample rate in Hz. Must match what the model was trained on. |
| channels | int | 1 | Number of channels. Always 1 (mono) — do not change. |
| dtype | str | float32 | Internal sounddevice format. Always float32. |
| record_duration_s | float | 30.0 | Recording window in seconds. Configurable per session. |
| device_index | int | None | ALSA/PortAudio device index. None = system default. Use list_audio_devices() to find USB mic index. |

## 4.2  Noise Reduction Parameters

| Parameter | Type | Default | Description |
| --- | --- | --- | --- |
| noise_estimate_duration_s | float | 0.5 | Seconds of initial audio treated as noise profile. |
| n_fft | int | 1024 | FFT window size in samples (64 ms at 16 kHz). |
| hop_length | int | 256 | Hop size in samples (16 ms). Controls overlap. |
| oversubtraction | float | 2.0 | Alpha factor. Increase for noisier environments (max ~4.0). |
| spectral_floor | float | 0.02 | Beta floor. Increase if you hear musical noise artefacts. |

## 4.3  Output and Detection Parameters

| Parameter | Type | Default | Description |
| --- | --- | --- | --- |
| target_peak_db | float | -3.0 | Peak normalisation target in dBFS. |
| clipping_threshold | float | 0.99 | Samples above this fraction of full scale are flagged as clipped. |
| clipping_warn_ratio | float | 0.001 | Warning threshold: warn if more than 0.1% of samples are clipped. |
| output_dir | str | "." | Directory to write output WAV files and summary. |
| output_prefix | str | "recording" | Filename prefix for output files. |

# 5. Model Input Contract

The following describes the exact format that audio_pipeline.py produces. The teammate responsible for model training must confirm that their model expects these values before integration.

| Property | Value | Notes |
| --- | --- | --- |
| Array shape | (num_samples,) | 1D numpy array — flat, no channel dimension |
| dtype | float32 | Native Python float32 — no conversion needed before inference |
| Sample rate | 16000 Hz | Must match the sample rate the model was trained on |
| Channels | 1 (mono) | Always mono — USB mic is configured single-channel |
| Amplitude range | [-1.0, 1.0] | Clipped to this range after normalisation |
| Peak level | -3 dBFS (configurable) | Set target_peak_db in PipelineConfig to change |
| Silence | Preserved | Pauses and silence are NOT trimmed — VAD is disabled by default |

|  | Integration handoff The PipelineResult.cleaned_audio field is the array to pass to inference. Access it as: result = run_pipeline(audio, config) model_input = result.cleaned_audio # shape (N,) float32, 16 kHz |
| --- | --- |

# 6. test_pipeline.py — Usage Guide

test_pipeline.py is a fully self-contained script. Run it directly on the Pi to record, process, and validate audio quality. It saves three files per run: raw WAV, cleaned WAV, and a text summary.

## 6.1  Quick-Start Commands

### Find your USB microphone device index

python test_pipeline.py --list-devices

Look for 'HyperX' or 'USB Audio' in the output. Note the card number and use it as --device.

### Standard 30-second recording

python test_pipeline.py --device 3

Saves raw and cleaned WAVs plus a summary to ./pipeline_output/. Filename prefix is timestamp-based (e.g. 20260425_121218_raw.wav).

### Custom duration

python test_pipeline.py --device 3 --duration 60

### Process an existing WAV file (skip recording)

python test_pipeline.py --input 20260425_121218_raw.wav

Useful for re-processing an earlier recording with different noise parameters without re-recording.

### Adjust noise reduction aggressiveness

# For a quiet indoor room (default): python test_pipeline.py --device 3 --noise-alpha 2.0 # For a noisy environment: python test_pipeline.py --device 3 --noise-alpha 3.5

### Custom output directory and prefix

python test_pipeline.py --device 3 --output-dir /home/raouls/recordings --prefix patient_001

### Debug: run without noise reduction

python test_pipeline.py --device 3 --skip-denoise

Useful for isolating whether an artefact comes from the filter or the noise reduction step.

### Debug: run without bandpass filter

python test_pipeline.py --device 3 --skip-filter

### Enable VAD trimming (diagnostic use only)

python test_pipeline.py --device 3 --skip-vad

|  | Do NOT enable VAD trimming for recordings intended for dementia screening inference. Pauses and silence are clinically significant — trimming them degrades model accuracy. Use --skip-vad only for diagnostic purposes such as checking if the subject spoke at all. |
| --- | --- |

### Verbose debug logging

python test_pipeline.py --device 3 -v

# 7. CLI Arguments — Full Reference

| Argument | Type | Default | Description |
| --- | --- | --- | --- |
| --duration | float | 30.0 | Recording window in seconds |
| --input | path | None | Process an existing WAV instead of recording |
| --device | int | System default | ALSA device index for the USB mic |
| --list-devices | flag | — | Print all audio devices and exit |
| --output-dir | path | ./pipeline_output | Directory for output files |
| --prefix | str | Timestamp | Filename prefix for output files |
| --noise-alpha | float | 2.0 | Spectral subtraction aggressiveness (1.0–4.0) |
| --noise-estimate-s | float | 0.5 | Seconds to use as noise profile from start of recording |
| --vad-threshold | float | -35.0 | VAD energy threshold in dB relative to peak |
| --skip-vad | flag | — | Re-enable VAD trimming (disabled by default) |
| --skip-denoise | flag | — | Skip spectral subtraction stage |
| --skip-filter | flag | — | Skip bandpass filter stage |
| -v / --verbose | flag | — | Enable DEBUG-level logging |

# 8. Output Files

| File | Format | Description |
| --- | --- | --- |
| {prefix}_raw.wav | 16-bit PCM WAV, 16 kHz mono | Unprocessed recording straight from the microphone |
| {prefix}_cleaned.wav | 16-bit PCM WAV, 16 kHz mono | Fully processed audio — ready to pass to the model |
| {prefix}_summary.txt | Plain text | Processing stats: durations, SNR before/after, clipping, config |

## 8.1  Summary File Contents

The summary file printed after every run includes:

- Total raw recording duration

- Speech duration after VAD (if VAD was enabled)

- Estimated SNR before and after noise reduction

- Clipping detection result and percentage of clipped samples

- Total processing time on-device

- List of stages applied

- Full config snapshot (bandpass range, noise alpha/beta, VAD threshold, normalisation target)

## 8.2  Listening to Output Files on Raspberry Pi 4

To compare raw vs cleaned audio directly on the Pi using the HyperX headset (card 3):

# Play raw aplay -D hw:3,0 -f S16_LE -r 16000 -c 1 pipeline_output/20260425_121218_raw.wav # Play cleaned aplay -D hw:3,0 -f S16_LE -r 16000 -c 1 pipeline_output/20260425_121218_cleaned.wav

To transfer files to a laptop for easier listening:

# Run this on your laptop: scp raouls@<PI_IP>:~/Final_Project/pipeline_output/*.wav ~/Desktop/

# 9. Resource Usage on Raspberry Pi 4

| Resource | Value | Notes |
| --- | --- | --- |
| Audio buffer (30 s) | ~1.9 MB | float32 at 16 kHz |
| FFT workspace | ~8 KB | 1024-point complex buffer |
| Spectral subtraction | ~3 MB peak | Overlap-add reconstruction buffers |
| Total pipeline peak RAM | ~8–10 MB | Well within Pi 4 limits |
| Processing time (30 s) | ~1.5–3 s | All four stages combined |
| Bandpass filter | ~50 ms | 5th-order sosfiltfilt |
| Spectral subtraction | ~200–400 ms | FFT/IFFT per frame with overlap-add |
| VAD (if enabled) | ~60–100 ms | Energy + ZCR per 30 ms frame |
| Normalisation | <10 ms | Single pass over array |

# 10. Installation & Dependencies

## 10.1  System Dependencies

# Install PortAudio (required by sounddevice) sudo apt update sudo apt install libportaudio2 # Optional: sox for richer audio playback with amplitude meter sudo apt install sox libsox-fmt-all

## 10.2  Python Dependencies

pip install sounddevice numpy scipy

No other packages are required. noisereduce and librosa are explicitly NOT used — they add ~200 MB and slow import time on Pi 4.

| Package | Version tested | Purpose |
| --- | --- | --- |
| sounddevice | 0.5.5+ | USB microphone capture via PortAudio |
| numpy | 2.4+ | Array operations and FFT computation |
| scipy | 1.17+ | Butterworth filter design and WAV file I/O |

# 11. Error Handling Reference

| Error / Situation | What happens | How to fix |
| --- | --- | --- |
| Microphone disconnected | MicDisconnectedError raised with clear message | Check USB connection: lsusb ¦ grep -i audio. Run --list-devices to verify index. |
| Wrong device index | MicDisconnectedError: device has no input channels | Run --list-devices and select the correct index for the USB mic. |
| Near-silent recording | Warning logged: max amplitude < 1e-7. Pipeline continues. | Check mic gain: alsamixer. Ensure subject is close to the mic. |
| Clipping detected | Warning logged with clipped sample count and percentage. | Reduce microphone gain in alsamixer. Move subject slightly away from mic. |
| No speech in recording (VAD enabled) | SilenceOnlyError raised. Raw WAV saved for debugging. | Lower --vad-threshold (e.g. -45). Check mic gain. Ensure subject speaks. |
| Wrong sample rate in input file | Warning logged. scipy.signal.resample applied automatically. | Have teammate export test files at 16 kHz to avoid resampling. |
| PortAudio library not found | OSError on import of sounddevice | Run: sudo apt install libportaudio2 |

# 12. Design Decisions

## sounddevice vs PyAudio

sounddevice was chosen over PyAudio. Both wrap PortAudio underneath so latency is identical, but sounddevice has a cleaner blocking API, no known memory leaks on Linux ARM, and its error handling maps cleanly to Python exceptions. PyAudio's callback interface requires manual buffer management and has documented instability on Raspberry Pi OS.

## Energy-based VAD vs webrtcvad vs silero-vad

An enhanced energy-based VAD (RMS + zero-crossing rate + hangover) was implemented rather than webrtcvad or silero-vad. webrtcvad is excellent but is a C extension that can have installation issues on ARM and produces only binary frame labels with no confidence score. silero-vad is a neural model (~2 MB ONNX) that requires PyTorch or ONNX Runtime and adds 15–30 ms per frame on Pi 4 — significant overhead for a pre-processing step. The energy-based approach runs in < 1 ms per frame, has no extra dependencies, and for controlled indoor recordings achieves 85–90% of webrtcvad accuracy. The VAD interface is designed to be swappable — if edge cases arise, webrtcvad is a drop-in replacement.

## scipy spectral subtraction vs noisereduce

scipy-based spectral subtraction was implemented rather than the noisereduce library. noisereduce uses spectral gating which produces slightly cleaner output, but it depends on librosa which adds approximately 200 MB of dependencies and significantly slows Python startup on Pi 4. The scipy implementation delivers 80–90% of the quality with a ~3 MB memory footprint and no additional packages.

## Why pauses are preserved

VAD trimming is disabled by default because long pauses, hesitations, and silence are primary acoustic markers for early cognitive decline as identified in the project proposal. Trimming silence would destroy the temporal structure the model relies on to detect these patterns. Noise reduction and normalisation are still applied to improve signal quality without altering timing.

## Lazy sounddevice import

sounddevice is imported lazily (only when capture_audio() is called) rather than at module load time. This allows audio_pipeline.py to be imported in environments without PortAudio installed — for example, on a development machine or CI system where only the signal processing functions are needed.

# 13. Programmatic Usage

To use the pipeline from another Python module (e.g. the inference wrapper), import and call run_pipeline() directly rather than subprocess-calling test_pipeline.py.

## 13.1  Minimal integration example

from audio_pipeline import PipelineConfig, run_pipeline, capture_audio config = PipelineConfig( record_duration_s=30.0, device_index=3, # HyperX headset on Pi 4 oversubtraction=2.0, ) # 1. Capture raw_audio = capture_audio(config) # 2. Process result = run_pipeline(raw_audio, config) # 3. Pass to model model_input = result.cleaned_audio # float32, shape (N,), 16 kHz # risk_score = model.infer(model_input)

## 13.2  Processing an existing file

from audio_pipeline import PipelineConfig, run_pipeline, load_wav, save_wav audio, sr = load_wav('recording.wav') config = PipelineConfig(record_duration_s=len(audio)/sr) result = run_pipeline(audio, config) save_wav(result.cleaned_audio, 'cleaned.wav')

## 13.3  Skipping stages

from audio_pipeline import PipelineConfig, run_pipeline, PipelineStage # Run only the bandpass filter and normalisation result = run_pipeline( audio, config, skip_stages={PipelineStage.DENOISED, PipelineStage.VAD_TRIMMED} )

# 14. Ethical and Regulatory Considerations

|  | This system is a screening support tool only. It is NOT a certified medical device and must NOT be used as the sole basis for any clinical decision. All outputs must be reviewed by a qualified healthcare professional. |
| --- | --- |

- All audio recordings must be collected with explicit informed consent from the subject.

- Recordings should be stored securely and not transmitted over the network without encryption.

- The system produces risk indicators — false positives and false negatives are expected and must be communicated to clinicians.

- No personally identifiable information should be embedded in filenames or stored alongside audio files without appropriate data governance controls.

- Validation by Dr. Khaled Sidani (Resident Neurosurgeon, AUBMC) is noted in the project proposal — clinical validation of the full pipeline including the model is required before any clinical deployment.

Edge AI Dementia Voice Detector — Audio Pipeline Documentation — Confidential